AUDITOR NAME: _____________________________________      AUDIT DATE: _____________

LOCATION/UNIT: _____________________________________

autoclave make and model (optional): ____________________________________

| Element Assessed | Response | Notes |
| --- | --- | --- |
| Employee Competency |  |  |
| 1. All employees performing autoclave sterilization at the site have completed a competency? | Yes No |  |
| - Competency is on file? | Yes No |  |
| Clean and Soiled Rooms |  |  |
| 3. Are there separate decontamination and clean rooms? | Yes No |  |
| 4. Is the autoclave in the clean room? | Yes No NA, all one room |  |
| - Does the room have unidirectional flow from contaminated to clean? | Yes No |  |
| - Supplies necessary for adherence to hand hygiene are readily accessible to healthcare personnel in areas where reprocessing occurs: |  |  |
| - Soap, water, and paper towels | Yes No |  |
| - Alcohol-based hand rub | Yes No |  |
| - Are decontamination and clean rooms free of dust, dirt, soil, and clutter? NOTE: Specify concerns. | Yes No |  |
| - Are ceilings/ceiling tiles intact? | Yes No |  |
| 9. Are ceilings, floors, or walls free from water damage or mold? | Yes No |  |
| Instrument Cleaning and Inspection |  |  |
| 10. Soiled items are pre-cleaned at the point of use prior to transport? | Yes No Unable to observe |  |
| 11. Soiled items are kept moist for transport (e.g. foam, gel, cleaning solution)? | Yes No Unable to observe |  |
| 12. Contaminated items are properly contained during transport? | Yes No Unable to observe |  |
| 13. Employee(s) correctly uses personal protective equipment? Specify concerns in the notes section. | Yes No Unable to observe |  |
| 14. For manual cleaning, is fresh enzymatic solution prepared for each use and according to the manufacturer’s recommendations? NOTE: Review quality controls for ultrasonic cleaner and automated washer chemical dispensing. | Yes No N/A, facility uses only mechanical cleaning. |  |
| 15. Are manufacturer’s instructions for cleaning, disinfection/sterilization immediately available? | Yes No |  |
| 16. Are items cleaned and lubricated (if indicated) properly? - Item soaked in the open position for at least 1-5 minutes - Instruments are brushed below the level of the enzymatic solution - All surfaces come in contact with the cleaning solution - Rinsed in warm tap water - Sprayed with a lubricant or dipped in lubricant (optional – may also occur in a washer/disinfector cycle) - Air dried on a towel or chux in the open position | Yes No N/A, facility uses only mechanical cleaning. |  |
| 17. Brushes should either be single use, disposable items or if reusable should be inspected for wear, cleaned after each use and disinfected or sterilized at least once a day. | Yes No N/A, facility uses only mechanical cleaning. |  |
| 18. For mechanical cleaning, are quality controls completed for ultrasonic and washer disinfectors/washer sterilizers? | Yes No N/A, facility uses only manual cleaning. |  |

| 19. Are instruments inspected after cleaning? If yes, what is verified? | Yes, for adequate cleaning Yes, for integrity and functionality Yes, for cleaning, integrity and functionality Not properly inspected |  |
| --- | --- | --- |
| Instrument Pouching and Wrapping |  |  |
| 20. Are instruments pouched appropriately? - Pouch appropriate size for instruments - Instruments are dry - Sharp instruments are placed in tip protectors - Instruments are in the open unlocked position - Handles are closest to the peel area of the pouch - Chemical indicator/ integrator is in pouch and visible - Pouch is closed at the pre-fold line/perforation without any gaps - If double pouching is utilized, the inner pouch is smaller than the outer pouch and not folded. Pouches are paper to paper and plastic to plastic. | Yes No | Indicate the number of pouches assessed: Indicate the number of pouches that were correct: |
| 21. Are pouches and/or packs labeled and stored appropriately? - One step wrap or double layer wrap is appropriate in size and not cut - Wrap secured with a maximum of three strips of steam indicator tape - Alcohol based marker or gun label used - Labeling on the plastic side of the pouch - Wrapped packages: Label on outside wrap - Include: date sterilized. Load number, initials, department if shared, sterilizer name/number if more than one autoclave - Pouch/wrap integrity maintained | Yes No |  |
| 22. Are instrument trays prepared properly? - Curved tips are positioned in the same direction - Cupped or concaved instruments are positioned to avoid water/condensation collection - Chemical indicator/ integrator is in the tray | Yes No Unable to observe |  |
| 23. If a rigid sterilization container system is used, are manufacturer’s written instructions for use regarding set preparation and assembly available? NOTE: Rigid container systems and medical devices must be tested and validated for preset standard steam sterilization cycles. See AAMI 13.10 Periodic product quality assurance testing of rigid sterilization container systems for additional details. | Yes No NA, a rigid container system is not used |  |
| Autoclave |  |  |
| 24. Is current autoclave policy and procedure available? NOTE: If no, skip to question 25. | Yes No |  |

| 25. Has the policy/procedure been reviewed in the last year? | Yes No Unknown |  |
| --- | --- | --- |
| 26. Are manufacturer’s instructions for the autoclave readily accessible? | Yes No |  |
| 27. Is there documentation that the autoclave has undergone preventive maintenance in the last year? | Yes No |  |
| 28. Is documentation on the autoclave maintenance log maintained and up to date? (Routine maintenance) | Yes No |  |
| 29. Are pouches and wrapped packs appropriately placed in the autoclave? - Pouches: Lying flat with paper side down in a single layer or standing on edge with correct tray accessory all facing the same side - Wrapped Packs: Solid trays positioned on edge or perpendicular. Perforated trays loaded with tray bottom down. | Yes No Unable to observe |  |
| 30. Are chemical indicators read and documented after pouch/tray has gone through the sterilization cycle? NOTE: Rejected pouches shall be re-pouched with a new chemical indicator and run in the next autoclave load. | Yes No |  |
| 31. Is biological indicator testing performed preferably daily, but at least weekly and with each load for implants? | Yes No |  |

| 32. After sterilization, are pouches or packs stored in a designated clean area so that sterility is not compromised? - 8-10 inches from the floor - 18-20 inches from the ceiling - 2 inches from an outside wall - Closed shelving/storage is preferred - “First in, First out” stock rotation | Yes No |  |
| --- | --- | --- |
| Recall of Sterilized Items due to a Positive Biological Indicator |  |  |
| 33. Is there a recall policy ad procedure available? NOTE: It should include the process to track sterilized instruments used on patents. If response is no, tool is complete. | Yes No |  |
| 34. Has the recall policy been reviewed in the last year? | Yes No Unknown |  |